## Slide 1: K-SYNC
OPTIMIZING THE CONDUCTOR OF HYDRO-ACOUSTICS

## Slide 2: System Overview
1

## Slide 3: Course Objectives
Understand K-Sync Capabilities
Identify and describe K-Sync components and integration on vessel
Perform startup, shutdown, and basic operation of K-Sync

## Slide 4: System Overview
K-Sync controls pinging of each echosounder with real-time trigger display.
Avoids interference from echosounders with similar frequencies or harmonics that will cause bad data collection.
K-Sync receives ready signals from echosounders
Supports up to 16 groups in real time or manual depth input
Trigger from external input, calculated or fixed time
Notes: Idea*** - your conductor for the myriad of systems so that interference is limited. Symphony of sound!
From our amazing website:
The KONGSBERG synchronizing unit, K-SYNC is designed to solve these problems by proper timing of the instruments and by controlling the triggering of each instrument's transmission.

## Slide 5: Frequencies and Interference
The Problem
EQUIPMENT | FREQUENCY
Sub Bottom Profiler | 0.5-30 kHz
HiPAP | 20-30 kHz
EM124 | 10.5-13.5 kHz
EM304/MKII | 26-34/20-32 kHz
EM712 | 40-100 kHz
EM2040 | 200-400 kHz
EK80 | 18-333 kHz
ADCPs | 38 – 300 kHz
ME 70 | 70 kHz
Notes: Some sonars have overlapping ranges or harmonics resulting in interference between systems.

## Slide 6: Synchronization
Synchronize your systems together!
Note: Using a K-SYNC unit with several active acoustic systems will slow down ping rate and reduce the data density on the sensors involved.
The Solution
Workstation
PU
Acoustic system
Notes: The K-sync works to ping the systems in a way that prevents interference.
This is achieved through a combination of user settings and K-sync software intelligence

## Slide 7: Hardware and Connections
2

## Slide 8: System Integration
Hardware and Connections

## Slide 9: Processing Units (PU)
Hardware and Connections
19-Inch rack-mounted K-SYNC PU
1K-SYNC PU for wall-mount
Notes: Your system is Rack Mounted

## Slide 10: PU Internal Components
Hardware and Connections
A Power
B 16 modules for interfacing with different acoustic systems
19-Inch rack-mounted K-SYNC PU
1K-SYNC PU for wall-mount
Notes: Up to 16 module for acoustic connections

## Slide 11: Modules
There are 16 Modules to connect
A – Wire terminal connections
1 through 6 as indicated by the arrows
B – Configuration board
Hardware and Connections

## Slide 12: Module Connections
Hardware and Connections

## Slide 13: Software Overview
3

## Slide 14: Software Overview
Notes: Network based SU operator sw may be run on any PC
System setup
System monitoring
If setup the K-sync SU will continue to trigger the echosounders
- even if the operator sw is closed
CAUTION! Always power up the K-Sync before pinging echo sounders
Trigger outputs are undefined during power-up (first 20 seconds) and signal levels could potentially be interpreted as trigger pulse by the echo sounders until signal levels have settled
Systems that are currently connected to the K-sync in image
EM 124
EM 2040
EM 712
SBP 29
OS 150
OS 38
OS75
Workhorse 300
EK80
ME 70

## Slide 15: Synchronizing Unit Operation
Software Overview

## Slide 16: System Control
Software Overview
Notes: On the left side of the window is the System control, status, and User Configuration profiles
You can create and save setup configurations based on priority and project goals
These are user selectable when the time arises

## Slide 17: System Status
Software Overview
Depth is available
Override incoming depth
Incoming depth status

## Slide 18: System Status
Software Overview
Override incoming depth
Incoming depth status
Note: Sometimes manual depth is best when doing pelagic sampling in deep water.

## Slide 19: Activity Display
Software Overview
Notes: Sonar Activity is shown in the red box (aka real time trigger display)

## Slide 20: Realtime Trigger Display
Software Overview
Yellow line separates the trigger groups
OS38 is busy much longer than EM710 and determines the ping rate
Trigger
Transmit and busy
Ready
Note: The displayed pulses are indicators. They do NOT reflect active high/low settings
Notes: Blue jump = Ksync triggered a transmit pulse
Red jump = system transmitted
Red line = system is waiting for the ready signal ( still listening for return)
Green line = system is done pinging/ listening and waiting for next trigger

## Slide 21: Echo Sounder Status
Software Overview
Echosounders may be disabled
Disabled echosounders will not be triggered - even if they are member of a group

## Slide 22: Trigger Group
Software Overview
Active Group
Set up groups
Notes: Shows which group each echo sounder is assigned to
Trigger groups
Several (or one) echosounder in each group.
Specify the transmit sequence of each group.

## Slide 23: Trigger Group Configuration
Software Overview
If some echo sounders can ping several times before next instrument, make many equal groups

## Slide 24: Trigger Group Example
Software Overview
Empty groups are skipped
Active Group
Shows which group each echosounder is assigned to
In this case: EM122, EM710 and SBP120 ping 4 times between each Workhorse 300

## Slide 25: Diagnostic Tool
Software Overview
Shows signal level and pulses on each I/O
(Green = 5v(high)
Module & pin numbers refer to physical connections inside the K-Sync
Terminal ID refer to selections in the installation settings
Notes: Shows signal level and pulses on each input and output signal. (Green = 5v(high)) Module and pin numbers refer to physical connections inside the SU.
Terminal ID refer to selections in the installation settings.

## Slide 26: Installation Settings
Software Overview
To be able to change any installation settings:
Press “Allow modification” and input correct password
It is helpful to save a backup of all of these settings in a separate location
Notes: Password simrad0
How to backup: Copy the folder to sperate location:
C:\ProgramData\Kongsberg Maritime\K-Sync Application\config
System settings are stored as an XML file
“Program folder”…\config\system settings.xml
Factory settings are stored under
“Program folder”…\system\
Installation settings are in separate files
make.model.installation.xml

## Slide 27: System Settings
Software Overview
Settings -> System settings
Approximate sound speed: Used for travel time calculations
Sync unit’s IP address and port number.
Default Subnet mask: 255.255.240.0
Notes: Define K-SYNC IP address and ports
Define type of depth data gram
Set approx. sound speed
Supported depth datagram formats:
NMEA depth (DBS, input datagram)
NMEA DPT (input datagram)
EAx: Single beam depth (EA 500 input datagram)
EM Depth ‘D’: Multibeam depth (e.g. EM 120/1002/3002)
EM Depth ‘X’: Multibeam depth (e.g. EM 122/302/710/2040)
EM Depth ‘E’: Single beam depth (e.g. EA 600)

## Slide 28: Manage Echosounders
Software Overview
Password: simrad0
Manage Installed Echosounders Mark the echosounders that are physically connected.
Add new echosounder if it is not already in the list.
To add or remove echo sounders from the Installed Systems
Settings -> Installation -> Installation settings
Notes: Settings -> Installation -> Installation settings

## Slide 29: Add New Echosounder
Software Overview
Notes: Make: Manufacturers Name
Model: Sounder type

## Slide 30: Installation Settings
Software Overview
Notes: It is helpful to note which echosounder is connected to which module number
Connect each sounder to the appropriate module
For All signal types (Ready to transmit, Transmitting, and trigger output)
Mark Signal is available-- IF feedback signals from echosounders are available
  Specify terminal connection and Module for Each signal available
  Specify Signal Activity
    Is signal active high or low?
Define pulse properties
Hit Apply then OK when complete

## Slide 31: Runtime Settings
Software Overview
Settings -> Runtime settings
Calculated: Ping rate is based on depth and sound speed
External input: Automatic sequence calculations
Notes: Calculate
Multiplication factor: increase ping rate echosounders due to longer range than depths.
Fixed Delay: Adds a fixed delay / min. ping period.
Fixed period:
Ping rate is depth independent
Maximize pinging allows that system to ping more while waiting for other systems

## Slide 32: Optimizing Settings
4

## Slide 33: Knowing Crosstalk
Optimizing Settings
“Rainy” Echogram
Noise at certain depth
Notes: Calculate
Multiplication factor: increase ping rate echosounders due to longer range than depths.
Fixed Delay: Adds a fixed delay / min. ping period.
Fixed period:
Ping rate is depth independent
Maximize pinging allows that system to ping more while waiting for other systems