Tritium Laboratory

May 8, 2006

SWAB REPORT #408 - Addendum

SWAB DATE: 5 April 2006 (USAP Van #1)

5 May 2006 (Tritium Laboratory)

R/V Nathaniel B. Palmer

_________________________

Dr. James D. Happell

Research Associate Professor

Distribution:

SWAB Committee

Bob Kluckhohn

REPORT FOR SWAB # 408 - Addendum

LOCATION  : Punta Arenas, Chile & Miami, Florida DATE  : 4/5/06, 5/5/06

TECHNICIAN: Cecilia Roig                         STATUS: Van #1 has

VESSEL/LAB: USAP Van #1 & Tritium Lab benchtop   tritium contamination,

Tritium Lab benchtop did not show isotope spills.

SAMPLE  SAMPLE IDENTIFICATION               	 NET ACTIVITY EXTRACTED

#                                           	3H dpm/m2    14C dpm/m2

1   Machine Blank						     -	   -

USAP Van # 1

54   Inside fume hood                                  8,450*        0

55   Workbench left of sink                           15,971**       0

56   Workbench across fume hood                        4,565*        0

57   Inside Consul 230 freezer top                     4,737*        0

58   Inside Consul 230 refrigerator bottom            77,484**       0

59   Deck inside door                                 53,646**       6

60   Deck left of workbench                           11,506**       0

61   Deck in front of fumehood                        36,183**       0

62   Workbench right of sink                           2,468*        0

63   Drawer/basket inside Consul 230               1,054,535***    118

64   Final Bucket blank C.O. #2                            0         0

Tritium Laboratory

65   Initial Bucket blank                                  0         0

66   Initial laboratory benchtop swab                     30        16

67   Laboratory benchtop after use of Windex               0         0

68   Laboratory benchtop after use of Dow                116        17

69   Laboratory benchtop after use of both Windex and Dow 13         0

70   Final Bucket blank                                    0         0

COMMENTS

USAP had concerns regarding accuracy of numbers reported for samples collected on USAP Van #1 (SWAB Report #408).  All samples collected in USAP Van #1 were reprocessed using new LSC vials and recounted (each new vial was recounted fourteen times).  The results concur with numbers reported in SWAB Report #408.  The outcome shows the same areas in Van #1 contaminated, at the same levels, after running the second set of LSC vials.

The possibility of household cleaners (Windex and Dow) affecting the samples was also brought up.  Swab tests were done at the Tritium Laboratory using these cleaners.  Areas swabbed at the Tritium Laboratory with the cleaners were not contaminated.